## Slide 1: 2d lidar sensors TiM3xx-7xx plane alignment
8PM/83
2019

## Slide 2: TiM3xx-7xx
Tolerance specification:
Our official specification is a kind of worst case scenario.
It is defined by the tolerance chain due to mounting of the mechanical components
Mounting
Mounting has a strong influence on the plane alignment.
Therefore we recommend using a mounting bracket with inbuilt adjustment. E.g.: https://www.sick.com/us/en/mounting-kit/p/p492344
Tolerance summary from the housing
Roll | Pitch | Yaw
 3° |  | 1° for TiM551 0.33° for TiM571

## Slide 3: TiM3xx-7xx
Considering spot size
Considering spot size & laser divergence (9.5 mrad = ~0.544°)
Modeling the laser emitter
Spot size @ screen = 6 mm
9.5/2 mrad = ~0.272°
Spot size + divergence of 9.5mrad @ 10m: ~101mm
Spot size + divergence of 9.5mrad @ 250mm: ~8.4mm
Spot size + divergence of 9.5mrad @ 500mm: ~10.8mm
9.5/2 mrad = ~0.272°

## Slide 4: TiM3xx-7xx
Considering scan field flatness 3°
Modeling the scan field flatness
3°
3°
Scan field flatness @ 10m: 10,000 mm x tan(3°)= ~524 mm
Scan field flatness @ 250mm: 250 mm x tan(3°)= 13,1 mm
Scan field flatness @ 500 mm: 500 mm x tan(3°)= ~ 26,2 mm

## Slide 5: TiM3xx-7xx
Considering scan field flatness + spot size + divergence
WORST CASE Scenario (e.g. for desiging an light output slot )
Modeling the scan field flatness + laser emitter
9.5 mrad = ~0.544°
3°
3°
@ 10,000 mm distance
Scan field flatness 10,000 mm x tan(3°)= ~524 mm + ½ Spot size 3 mm +divergence of ½ 9.5mrad:~47,5 mm
= ~574,5 mm
@ 250mm distance
Scan field flatness 250 mm x tan3°: ~13,1 mm
+ ½ Spot size @ screen : 3 mm
+ divergence of ½ 9.5mrad: ~1.2 mm
= ~ 17.3 mm
@ 500 mm distance
Scan field flatness 500 mm x tan 3°: ~26,2 mm + ½ Spot size: 3 mm + divergence of ½ 9.5mrad: ~2.4 mm
= ~ 31,6 mm
½ Spot size @ screen = 3 mm

## Slide 6: Possible Misalignments
  Cone (e.g. due to internal misalignment of components)
  Tilt plane (e.g. due to internal misalignment of components)
    This can be compensated with the alignment bracket (https://www.sick.com/us/en/mounting-kit/p/p492344)
Effects can be mixed
The preceding explanations and calculations may be expressed in the following two effects.

## Slide 7: Many thanks for your attention.
8PM/83